KOGGER
kogger.tech

Serial Binary Protocol (SBP) specification

Document Information
Number	KS_SBP_100
Date	Sep 8, 2025
Revision	4.0.7
Status	Early Production Information

# Introduction
## Protocol frame structure
Header		Range over which the checksum is calculated					Checksum
SYNC1	SYNC2	ROUTE	MODE	ID	LENGTH	PAYLOAD	CHECK1	CHECK2
U1	U1	U1	U1	U1	U1	BYTE[LENGTH]	U1	U1
0xBB	0x55	BITFIELD	BITFIELD	1 … 255	0 … 128	BYTEARRAY	0 … 255	0 … 255

ROUTE
Name	Bits	Description
DEV_ADDRESS	0:3 bit	Device address. Default and broadcast address is 0x0.
RESERVED	3:7 bit	Reserved

MODE
Name	Bits	Description
TYPE	0:1 bit	Field defines the type and purpose of the data / 0 — Reserved, / 1 — CONTENT: DEVICE → HOST / 2 — SETTING: HOST → DEVICE / 3 —GETTING: HOST → DEVICE
RESERVED	2 bit	Reserved
VERSION	3:5 bit	Field defines the payload data version
MARK	6 bit	Once device is switched on, this flag is always in reset state (ZERO). It can be set to active state (ONE) by the host (see the ID_MARK) and the slave device keeps the flag in active state in every frame until hardware reset occurs or is reset by the host. Therefore the host monitors the device's actual settings.
RESPONSE	7 bit	HOST → DEVICE: / Set the flag to active state (ONE) in order to get the result of processing the command. The flag doesn't affect the response if one is provided by the TYPE field. / DEVICE → HOST: / The flag is in reset state (ZERO) by default. Payload goes according to the command specification. If flag is set, the payload contains the result of command processing (see RESP).

## Checksum
The checksum algorithm used is the Fletcher-16.
Example source code for calculating the checksum:
uint8_t CHECK1 = 0;
uint8_t CHECK2 = 0;

void CheckSumUpdate(uint8_t byte) {
CHECK1 += byte;
CHECK2 += CHECK1;
}
## Number Formats
All multi-byte values are ordered in Little Endian format.
All floating point values are transmitted in IEEE754 single or double precision.
All bit-field in LSB format.
Name	Type	Size (Bytes)	Range
S1	int8_t	1	-128 ... 127
U1	uint8_t	1	0 … 255
S2	int16_t	2	-32768 … 32767
U2	uint16_t	2	0 … 65535
S4	int32_t	4	-2'147'483'648 ... 2'147'483'647
U4	uint32_t	4	0 … 4'294'967'295
F4	float	4	-1*2^+127 ... 2^+127
D8	double	8	-1*2^+1023 ... 2^+1023

## Confirmation key
KEY_CONFIRM = 0xC96B5D4A

# Command specification
## Command overview
Name	ID	Description
Measurement data
ID_TIMESTAMP	0x01	Timestamp
ID_DIST	0x02	Distance data
ID_CHART	0x03	Chart data in reflection patterns
ID_ATTITUDE	0x04	Attitude
ID_TEMP	0x05	Temperature data
Settings data
ID_DATASET	0x10	Dataset management for automatic output
ID_DIST_SETUP	0x11	Detection Settings to Get Distance
ID_CHART_SETUP	0x12	Chart Settings
ID_DSP	0x13
ID_TRANSC	0x14	Transceiver settings
ID_SOUND	0x15	Sound speed settings
ID_PIN	0x16	Pin functions settings
ID_BUS	0x17	Bus settings
ID_UART	0x18	UART settings
ID_I2C	0x19	I2C settings
ID_CAN	0x1A	CAN settings
ID_IMU_SETUP*	0x1B	IMU settings
System
ID_VERSION*	0x20	Software and hardware version information
ID_MARK	0x21	Setting the mark of continuous work (non-reboot) device
ID_DIAG*	0x22	Diagnostic data
ID_FLASH	0x23	Work with built-in non-volatile memory
ID_BOOT	0x24	Boot device
ID_UPATE	0x25	Firmware update
Navigation
ID_NAV	0x64

* In developing
## RESP
Contains the result of command processing. Can be used as a check if the command is processed correctly. The MODE and ID fields are the same as in the initiating command.
Message format
TYPE	Version	Direction		Description
CONTENT	ANY	DEVICE → HOST
Format: [U1, U1, U1]; Length: 3
		Type	Range	Default	Unit	Description
CODE		U1				The field contains the response code of processing the command. / RESP_NONE = 0, / RESP_OK = 1, / RESP_ERR_CHECKSUMM = 2, / RESP_ERR_PAYLOAD = 3, / RESP_ERR_ID = 4, / RESP_ERR_VERSION = 5, / RESP_ERR_TYPE = 6, / RESP_ERR_KEY = 7, / RESP_ERR_RUNTIME = 8
CHECK1		U1				Command checksum
CHECK2		U1

## ID_TIMESTAMP(ID 0x01)
Message format
TYPE	Version	Direction		Description
GETTING	0	HOST → DEVICE		Request Timestamp
Format: []; Length: 0
		Type	Range	Default	Unit	Description
No data

Message format
TYPE	Version	Direction		Description
CONTENT	0	DEVICE → HOST		Timestamp from Device
Format: [U4]; Length: 4
		Type	Range	Default	Unit	Description
TIMESTAMP		U4	0 …		ms	Timestamp

## ID_DIST (0x02)
Message format
TYPE	Version	Direction		Description
GETTING	0	HOST → DEVICE		Get Distance
Format: []; Length: 0
		Type	Range	Default	Unit	Description
No data

Message format
TYPE	Version	Direction		Description
CONTENT	0	DEVICE → HOST		Data of chart

		Type	Range	Default	Unit	Description
DISTANCE		U4	0 ...		mm	Distance

Message format
TYPE	Version	Direction		Description
CONTENT	1	DEVICE → HOST		Data of chart

		Type	Range	Default	Unit	Description
NUMBER		U1
STRONG		U1
DISTANCE		U4	0 ...		mm
WIDTH		U2	0 ...		mm

## ID_CHART (0x03)
Request CHART
TYPE	Version	Direction		Description
GETTING	0	HOST → DEVICE		Get data of CHART
Format: []; Length: 0
		Type	Range	Default	Unit	Description
No data

Chart data
TYPE	Version	Direction		Description
CONTENT	0	DEVICE → HOST		Data of chart in sample format
Format [U2, U2, U2, U1[N]]; Length (6 + N)
		Type	Range	Default	Unit	Description
SEQ_OFFSET		U2	0 …			Sample Offset in Sequence
SAMPEL_RESOL		U2	10 …		mm	Samples resolution
ABS_OFFSET		U2	0 ...			Absolute Offset in sample number for Sequence
CHART		U1[N]	ARRAY			Data of chart. The maximum data length in one packet is 250.

Chart data
TYPE	Version	Direction		Description
CONTENT	1	DEVICE → HOST		Data of chart in sample format
Format [U2, U2, U2, U1[N]]; Length (6 + N)
		Type	Range	Default	Unit	Description
SEQ_OFFSET		U2	0 …			Sample Offset in Sequence
SAMPEL_RESOL		U2	10 …		mm	Samples resolution
ABS_OFFSET		U2	0 ...			Absolute Offset in sample number for Sequence
CHART		U1[N]	ARRAY			Data of two charts. The two channels are represented by interleaved data. For example, the first byte is the first measurement of the first channel and the second is the first measurement of the second channel. The maximum data length in one packet is 250. Accordingly, the maximum channel length is 125.

## ID_ATTITUDE (0x04)
Attitude
TYPE	Version	Direction		Description
GETTING	0, 1, 2	HOST → DEVICE		Request attitude
Format: []; Length: 0
		Type	Range	Default	Unit	Description
No data

Attitude
TYPE	Version	Direction		Description
CONTENT	0	DEVICE → HOST		Attitude data in Euler 321 format
Format: [S2, S2, S2]; Length: 6
		Type	Range	Default	Unit	Description
YAW		S2	-18000 ...18000		0.01 deg	Yaw
PITCH		S2	-18000 ...18000		0.01 deg	Pitch
ROLL		S2	-18000 ...18000		0.01 deg	Roll

Attitude
TYPE	Version	Direction		Description
CONTENT	1	DEVICE → HOST		Attitude data in quaternion format
Format [F4, F4, F4, F4]; Length 16
		Type	Range	Default	Unit	Description
W0		F4			—
W1		F4			—
W2		F4			—
W3		F4			—

## ID_TEMP (0x05)
Request temperature
TYPE	Version	Direction		Description
GETTING	0	HOST → DEVICE		Request temperature from device
Format: []; Length: 0
		Type	Range	Default	Unit	Description
No data

Temperature data
TYPE	Version	Direction		Description
CONTENT	0	DEVICE → HOST
Format: [S2]; Length: 2
		Type	Range	Default	Unit	Description
TEMP		S2			0.01 °C

## ID_DATASET (0x10)
Request Data set
TYPE	Version	Direction		Description
GETTING	0	HOST → DEVICE		Request Data set
Format: [U1]; Length: 1
		Type	Range	Default	Unit	Description
CHANNEL_ID		U1	0 … 2	0		Channel ID. Set 0 for request all active CHANNEL

Data set
TYPE	Version	Direction		Description
SETTING	0	HOST → DEVICE
CONTENT	0	DEVICE → HOST
Format: [U1, U4, U4]; Length: 9
		Type	Range	Default	Unit	Description
CHANNEL_ID		U1	0 … 2	0		Channel ID. Set 0 for reset all channel.
CHANNEL_PERIOD		U4	0 ...	0	ms	0 — for disable periodic response / >0 — for periodic response with period by value [ms]
CHANNEL_MASK		U4	BITFIELD	0x00		bit0 — ID_DIST Ver. 0, / bit1 — ID_CHART Ver. 0, / bit2 — ID_ATTITUDE Ver. 0, / bit3 — ID_ATTITUDE Ver. 1, / bit4 — ID_TEMP Ver. 0, / bit5 — ID_TIMESTAMP Ver. 0, / bit6 — DIST_NMEA_SDDBT,

## ID_DIST_SETUP (0x11)
Message format
TYPE	Version	Direction		Description
CONTENT	0	DEVICE → HOST		Data of chart
Format: [U4, U4]; Length: 8
		Type	Range	Default	Unit	Description
START_OFFSET		U4	0 ...	0	mm
MAX_DIST		U4	0 ...	50000	mm

## ID_CHART_SETUP (0x12)
Message format
TYPE	Version	Direction		Description
GETTING	0	HOST → DEVICE		Get setting of Chart
Format: []; Length: 0
		Type	Range	Default	Unit	Description
No data

Message format
TYPE	Version	Direction		Description
SETTING	0	HOST → DEVICE
CONTENT	0	DEVICE → HOST
Format: [U2, U2, U2]; Length: 6
		Type	Range	Default	Unit	Description
SAMPLE_COUNT		U2	1 ... 5000		—	Sample count
SAMPLE_RESOL		U2	10 ... 1000		mm	Sample resolution
SAMPLE_OFFSET		U2	0 …			Absolute Offset in number of sample

## ID_TRANSC (0x14)
Message format
TYPE	Version	Direction		Description
GETTING	0	HOST → DEVICE
Format: []; Length: 0
		Type	Range	Default	Unit	Description
No data

Message format
TYPE	Version	Direction		Description
CONTENT	0	DEVICE → HOST
SETTING	0	HOST → DEVICE
Format: [U2, U1, U1]; Length: 4
		Type	Range	Default	Unit	Description
FREQ		U2		675	kHz	Frequency
PULSE		U1		10	COUNT
BOOST		U1		1	—

## ID_SND_SPD (0x15)
Message format
TYPE	Version	Direction		Description
GETTING	0	HOST → DEVICE
Format: []; Length: 0
		Type	Range	Default	Unit	Description
No data

Message format
TYPE	Version	Direction		Description
CONTENT	0	DEVICE → HOST
SETTING	0	HOST → DEVICE
Format: [U4]; Length: 4
		Type	Range	Default	Unit	Description
SOUND_SPEED		U4		1500000	mm/s	Sound speed

## ID_UART (0x18)
Message format
TYPE	Version	Direction		Description
GETTING	0, 1	HOST → DEVICE
Format: [U4, U1]; Length: 5
		Type	Range	Default	Unit	Description
KEY_CONFIRM		U4		0xC96B 5D4A	—
UART_ID		U1	1 ...	1	—

Message format
TYPE	Version	Direction		Description
CONTENT	0	DEVICE → HOST
SETTING	0	HOST → DEVICE
Format: [U4, U1, U4]; Length: 9
		Type	Range	Default	Unit	Description
KEY_CONFIRM		U4	—	0xC96B 5D4A	—
UART_ID		U1	1 ...	1	—
BAUDRATE		U4	9600 … 921600	115200	bps

Message format
TYPE	Version	Direction		Description
CONTENT	1	DEVICE → HOST
SETTING	1	HOST → DEVICE
Format: [U4, U1, U1]; Length: 6
		Type	Range	Default	Unit	Description
KEY_CONFIRM		U4	—	0xC96B 5D4A	—
UART_ID		U1	1 ...	1	—
DEV_ADDRESS		U1	0...15	0

## ID_IMU_SETUP (0x1B) (In developing)
Message format
TYPE	Version	Direction		Description
SETTING	0	HOST → DEVICE		Calibrate Gyroscope
SETTING	1	HOST → DEVICE		Calibrate Accelerometr
Format: [U4]; Length: 4
		Type	Range	Default	Unit	Description
KEY_CONFIRM		U4		0xC96B 5D4A	—

## ID_VERSION (0x20) (In developing)
Message format
TYPE	Version	Direction		Description
GETTING	0	HOST → DEVICE
Format: []; Length: 0
		Type	Range	Default	Unit	Description
No data

Message format
TYPE	Version	Direction		Description
CONTENT	0	DEVICE → HOST
Format: [U4, U4, U4, U4, U4, U1[12], U2]; Length: 34
		Type	Range	Default	Unit	Description
SW_BOOT_VER		U4		—	—
SW_FW_VER		U4		—	—
HW_VER		U4		—	—
HW_FTRS		U4		—	—
SERIAL_NBR		U4		—	—
PART_NBR		U1[12]		—	—
FACTORY_DATE		U2		—	—

## ID_MARK (0x21)
Message format
TYPE	Version	Direction		Description
SETTING	0	HOST → DEVICE
Format: [U4]; Length: 4
		Type	Range	Default	Unit	Description
KEY_CONFIRM		U4		0xC96B 5D4A	—

Message format
TYPE	Version	Direction		Description
GETTING	0	HOST → DEVICE
Format: []; Length: 0
		Type	Range	Default	Unit	Description
No data

Message format
TYPE	Version	Direction		Description
CONTENT	0	DEVICE → HOST
Format: [U1]; Length: 1
		Type	Range	Default	Unit	Description
MARK		U1			—

## ID_DIAG (0x22) (In developing)
Message format
TYPE	Version	Direction		Description
GETTING	0	HOST → DEVICE
Format: []; Length: 0
		Type	Range	Default	Unit	Description
No data

Message format
TYPE	Version	Direction		Description
CONTENT	0	DEVICE → HOST
Format: []; Length:
		Type	Range	Default	Unit	Description
UPTIME		U4			ms
TEMP_IMU		S2			0.01 °C
TEMP_CPU		S2			0.01 °C
TEMP_MIN		S2			0.01 °C
TEMP_MAX		S2			0.01 °C
SYS_VOLT		U2			mV
BOOST_VOLT		U2			mV
DET_VOLT		U2			mV
DET_NOISE		U2			mV
AGC_GATE_VOLT		U2			mV

## ID_FLASH (0x23)
Save settings
TYPE	Version	Direction		Description
SETTING	0	HOST → DEVICE		Save run settings to non-volatile memory
SETTING	1	HOST → DEVICE		Restore run settings from non-volatile memory
SETTING	2	HOST → DEVICE		Erase non-volatile memory
Format: [U4]; Length: 4
		Type	Range	Default	Unit	Description
KEY_CONFIRM		U4		0xC96B 5D4A	—

## ID_BOOT (0x24)
Message format
TYPE	Version	Direction		Description
SETTING	0	HOST → DEVICE		Reboot device
SETTING	1	HOST → DEVICE		Run FW (for boot-loader mode)
Format: [U4]; Length: 4
		Type	Range	Default	Unit	Description
KEY_CONFIRM		U4		0xC96B 5D4A	—

## ID_UPDATE (0x25)
Message format
TYPE	Version	Direction		Description
SETTING	0	HOST → DEVICE		Upload firmware update (for boot-loader mode)
Format: [U2, U1[N]]; Length: (2+N)
		Type	Range	Default	Unit	Description
NBR_PACKET		U2	1 …
UPDATE_DATA		U1[N]	ARRAY

## ID_NAV (0x64)
Message format
TYPE	Version	Direction		Description
GETTING	0	HOST → DEVICE
Format: []; Length: 0
		Type	Range	Default	Unit	Description
No data

Message format
TYPE	Version	Direction		Description
CONTENT	0	DEVICE → HOST
Format: [D8, D8, F4]; Length: 20
		Type	Range	Default	Unit	Description
LATITUDE		D8			deg	Latitude
LONGITUDE		D8			deg	Longitude
ACCURACY		F4			m	Accuracy

## ID_DVL_VEL (0x79, 121)
Message format
TYPE	Version	Direction		Description
CONTENT	2	DEVICE → HOST
Format: [U4, U4, F4, F4, F4, F4, F4, F4, F4, F4, F4, F4, F4, F4, F4, F4, F4]; Length: 68
		Type	Range	Default	Unit	Description
FLAGS		U4
TIMESTAMP		U4			ms
DELTA_TIME		F4			s
LATENCY		F4			s
VELOCITY_X		F4			m/s
VELOCITY_Y		F4			m/s
VELOCITY_Z		F4			m/s
VELOCITY_Z1		F4			m/s
VELOCITY_Z2		F4			m/s
UNCERTAINTY_X		F4			m/s
UNCERTAINTY_Y		F4			m/s
UNCERTAINTY_Z		F4			m/s
UNCERTAINTY_Z1		F4			m/s
UNCERTAINTY_Z2		F4			m/s
DISTANCE_Z		F4			m
DISTANCE_Z1		F4			m
DISTANCE_Z2		F4			m

## ID_SIGNAL_ENCODER (0x66, 102)
Data message format
TYPE	Version	Direction		Description
SETTING	0	HOST → DEVICE
CONTENT	0	DEVICE → HOST
Format: [U4, U2, U1]; Length: 7
		Type	Range	Default	Unit	Description
RESERVED1		U4
BIT_LENGTH		U2	>=0			zero if there is no data to send
DATA		U1				values: 0 (reserved for sync), 1, 2, 3, 4, 5, 6, 7, 8 (commands)

Data request message format
TYPE	Version	Direction		Description
GETTING	0	HOST → DEVICE
Format: []; Length: 0
		Type	Range	Default	Unit	Description
No data

## ID_SIGNAL_DECODER (0x67, 103)
Data message format
TYPE	Version	Direction		Description
CONTENT	0	DEVICE → HOST
Format: [U4, U2, U1]; Length: 7
		Type	Range	Default	Unit	Description
TIMESTAMP		U4			ms
CARRIER_US		S8	>=0		us	precision time solution in us
CARRIER_CNT		S8	>=0		-	precision time solution in carrier counter
SOURCE_LVL		F4	-INF, +INF, NAN		db
SOURCE_SNR		F4	-INF, +INF, NAN		db
AZIMUTH		F4	-180...180, NAN		deg
ELEVATION		F4	-90...90, NAN		deg
RESERVED1		U4
RESERVED2		U4
BIT_LENGTH		U2	>= 0			zero if no data decoded
DATA		U1				values: 0 (reserved for sync), 1, 2, 3, 4, 5, 6, 7, 8 (commands)

Data request message format
TYPE	Version	Direction		Description
GETTING	0	HOST → DEVICE
Format: []; Length: 0
		Type	Range	Default	Unit	Description
No data